# Preface

The purpose of this document is to explain the procedures for QA of 3D printed cylinders and their associated reference sheet. This document assumes that a cylinder has been printed and processed. The overall workflow for 3D printed cylinders is: EUA (MRI/CT) -> Import and Registration -> Contours -> Preplan -> Export and 3D Model -> Print -> QA -> Sterilize -> Treat

Once the cylinders are printed, the Machinist (or a covering PA or physicist) will post-process the cylinders, including washing, curing, and tapping holes. Once this step is complete, the cylinder is ready for QA. The machinist will notify the brachy QA team that the cylinders are ready for QA.

# Cylinder QA Checklist

- Patients initials on cylinder (printed or etched)

- BBBBBBBB Sticker on cylinder

- Inspect cylinder for any cracks or damage

- Cylinder diameter is the same as EUA cylinder (verify in with Preplan)

- Channel holes pattern at base of cylinder are correct (compare to preplan numbering and orientation)

- Needle channels pass smoothly (with a stylette if plastic needle), no strenuous force needed

- Needle threads lock with correct needle (metal or plastic)

- Tandem passes smoothly with red cover

- Tandem thread locks

- Needles exit at expected trajectories (using brachyvision, preform, or brachify)

- Needles are not too close to each other (easy to lock collete with tool when nearby colletes are installed)

- Delivered cylinders to Brachy RTs to be taken for sterilization

# Reference Sheet QA Checklist

Reference sheet is exported to the patient folder:

- Base map is correct (compare to preplan numbering and orientation)

- Channel numbering is correct

- Tandem is channel 5

- Interstitial lengths are correct (brachify can include dead space, otherwise add 0.6cm for plastic, 1.0cm for metal needles)

- Sheet modified as necessary (open PDF with Word)

- Uploaded sheet to ARIA

Author: Michael Kudla

| Date | Revised by | Description |
| --- | --- | --- |
| 15-Oct-25 | MK | init |